## Slide 1: Gamma Ray and Anti-Matter Survey (GRAMS)
Power Distribution Unit (PDU)
Time Projection Chamber (TPC) SiPM Power & Telemetry Review

## Slide 2: Agenda
The Power Distribution Unit (PDU) for GRAMS is extensive
  Only reviewing the Silicon Photo-Multiplier (SiPM) PDU circuitry today
Overview of Requirements for SiPM bias, command & control
Block Diagram
Schematic Review
Grounding Configuration
Layout Review

## Slide 3: Requirements for TPC SiPM Power
The bias to the SiPMs shall be adjustable
  VMIN 43V  VMAX 60V
  Bias shall also have Enable / Disable Circuitry
  Telemetry Points to measure the Bias and Current
  The Current Draw for each SiPM is expected to be 20mA
  Temperature Range: -55C to +70C
    Target maximum altitude: 36km
Hamamatsu: S13370 Series

## Slide 4: Detector Information and Energy Ranges
GRAMS is using a Liquid Argon Time Projection Chamber (LArTPC) to detect the high energy events
  High energy particles through the argon will generate scintillated light to be collected by the SiPMs.
  Typical bubble chamber operation
    Superheated liquid is in the chamber and held just below its boiling point
    Interaction with the high energy particles will cause the liquid to boil along its path, creating ‘bubbles’. 
Generic bubble chamber image

## Slide 5: SiPM DC-DC Converter Choice
Ultravolt 1/16A12-P4 DC-DC Converter
  Able to supply 62V @ 4W.
    60V @ 20mA: 1.2W
  Has remote Enable / Disable Feature
  Has remote adjust feature for the output voltage

## Slide 6: 1/16A12-P4 Specifications
Applying +12V at the input.
Expecting a current draw of 20mA for each SiPM.
Output Bias will range from 40 to 60V.
Operational Temperature specs assume maximum load conditions
Input Parameters of the 1/16A12-P4 Converter |  |  | 
Parameter | Conditions | Values | Units
Input Voltage Range | Full Power | 11 to 16 | Volts
Current | Standby / Disabled | < 30 | mA
Current | Max Load @ output | < 100 | mA
AC Ripple | Max Load @ output | < 80 | mA(peak-peak)
Input Z | Nominal Conditions | 1.1 | MW
Enable/Disable Lvl | Nominal Conditions | Disable: 0 to 0.5 | Volts
 |  | Enable: < 2.4 | 
Output Parameters of the 1/16A12-P4 Converter (Vin = 12V) |  |  | 
Parameter | Conditions | Values | Units
Output Power | Max Load @ Output | 4 | Watts
Max Current Output | Entire Output Range | 64 | mA
Output V Ripple | 0.02% of Output V | 1.28 | mV
Temperature Range |  |  | 
Parameter | Maximum (Deg C) | Minimum (Deg C) | Range (Deg C)
Op. (@ Max Load) | -40 | 65 | 105
Storage | -55 | 105 | 160

## Slide 7: PDQ10-Q24-S5-D DC-DC Converter
Used these on COM-PAIR extensively to great success
Input Parameters of the PDQ10-Q24-S5-D Converter |  |  | 
Parameter | Conditions | Values | Units
Input Voltage Range | Full Power | 9 to 36 | Volts
Current | Maximum Input | 1.4 | A
Current | Max Load @ Input | 1.02 | A
Output Parameters of the 1/16A12-P4 Converter (Vin = 12V) |  |  | 
Parameter | Conditions | Values | Units
Output Power | Max Load @ Output | 10 | Watts
Max Current Output | Entire Output Range | 835 | mA
Output V Ripple | 0.02% of Output V | 100 | mV
Temperature Range |  |  | 
Parameter | Maximum (Deg C) | Minimum (Deg C) | Range (Deg C)
Op. (@ Max Load) | -40 | 105 | 145
Storage | -55 | 125 | 180

## Slide 8: Block Diagram of TPC SiPM Box
This box houses 10 boards
  One Telemetry Power Board
  Nine SiPM Power Boards
Telemetry Board
  Provides +12V to each SiPM DC-DC Converter
  Provides +5V that drives Telemetry on SiPM Boards
S11
S33
+12VIN
+12VRTN
=
=
+5V
+5VRTN
9 SiPM PDU Boards

## Slide 9: SiPM Power Board
Board Dimensions: 7” x 2.8”

## Slide 10: TPC SiPM Telemetry Power Board
Board Dimensions: 5.9” x 4.9”
Wurth Electronik Phoenix Socket: 1803471
Wurth Electronik Phoenix Contact: 1803620

## Slide 11: Schematic Review of SiPM Bias Board

## Slide 12: Input EMI/EMC Network

## Slide 13: Impedance and Cutoff Frequency of Input Filter
Some Calculations:
Impedances
Frequency

## Slide 14: Enable / Disable Circuitry
1/16A12-P4 Enable Signal
  There is a 10kW Pull-Up resistor internal to the Converter that defaults the block to enabled if no external control circuitry is added.
  Am using an opto-coupler to control this function
  ISP281D:
    Current Transfer Ratio (CTR): 300 to 600%
    ID = 100nA
    ISAT = 60 to 120mA
Truth Table |  | 
HV_EN | Low | High
Enable Line | Low | High
DC-DC Conv. | Disabled | Enabled

## Slide 15: Output Bias Control
Output Bias Control is from external signal
  Buffered in (LTC2052HV)
  R10 ensures bias is low if Bias_Ctrl is floating
Bias Control Levels |  | 
Bias_Ctrl | 0V | 5V
Output Voltage | 0V | 65V

## Slide 16: Output Bias Current Monitoring
The 1/16A12-P4 device has internal current monitoring
  Pin 3, I_OUT_MON
  IOUT = 20mA, I_OUT_MON = 19.7mA

## Slide 17: Output Bias Voltage Monitoring
Using a voltage divider and buffering output
  1MW resistor, (MBB0207VD1004BC100) is capable of dissipating 400mW.
  Vdrop across this resistor is: 57V is the output bias is set to 60V.
    PD = 3.3mW
  Vdiv = 0.05(Bias Output)
    60V will produce 3V at the input of U2B

## Slide 18: LED Indicator
Part Num: SSI-LXH072GD
Mounted on Chassis
1-1123824-1
+3V @ Bias Level = 60V
+2V @ Bias Level = 40V

## Slide 19: Input Current and Power Calculations
1/16A12-P4
  The expected load from each SiPM is approximately 20mA at maximum.
    Sub-system testing is on-going.
    VMIN 43V  VMAX 60V
    Manufacturer quotes efficiency at 90%
There are 9 boards pulling this
  11.97WTotal
  0.998A Total

## Slide 20: Telemetry I/O
All of the telemetry and control functions are sent and received through J3 (9-Pos Socketed D-SUB)
  Enable function is received as a differential signal (LVDS)
  U3, SN65LVDS2DBVR receives this signal, delivers to input of optocoupler (Q1)
  SEC_PWR_TELEM_RTN is tied to the secondary 5V_RTN on the board.

## Slide 21: Grounding Configuration
SiPM Telemetry Power Board
=
=
+12VIN
+12VRTN
+5VIN
+5VRTN
J2
J1
SiPM Power Board
=
=
V12V_RTN
SiPMBias_RTN
+12VIN
+12VRTN
49.9k
SiPMBias
SiPMBias_RTN

## Slide 22: Layout Review of SiPM Bias Board

## Slide 23: Layer Definition and Stack-Up
6 – Layer Board
  Layer 1: Routing
  Layer 2: Secondary Power RTN
  Layer 3: Power Rails (+5V, +3.3V)
  Layer 4: Routing
  Layer 5: Secondary Power RTN
  Layer 6: Chassis GND
Board Fab House: Sunstone Circuits
  Standard 6-Layer Stack up
Not expecting high currents or any high-speed signals

## Slide 24: Schematic Review of SiPM Telemetry Power Board

## Slide 25: Schematic of DC-DC Converter
PDQ10-Q24-S5-D is manufactured by CUI
  Used these extensively on COM-PAIR. No issues were seen.

## Slide 26: Impedance and Cutoff Frequency of Input Filter
Some Calculations:
Impedances
Frequency

## Slide 27: LED Power Indicator
Useful to have a visual indicator if the board is working as designed.
  LED will turn on if the +5V is being generated.

## Slide 28: Input Current and Power Calculations
PDQ10-Q24-S5-D
  The load from a single SiPM HV Board is ~ 52mA
    There are 9 SiPM Boards
    Total Current is ~ 470mA
Combine this with SiPM HV Board Current:
  Total Current Draw: 1.243A Total , 12V
  14.92Watts

## Slide 29: Layout Review of SiPM Telemetry Power Board

## Slide 30: Layer Definition and Stack-Up
6 – Layer Board
  Layer 1: Routing
  Layer 2: Secondary Power RTN
  Layer 3: Secondary Power Rail (+12VIN)
  Layer 4: Secondary Power Rail (+5VRTN)
  Layer 5: Secondary Power RTN
  Layer 6: Chassis GND
Board Fab House: Sunstone Circuits
  Standard 6-Layer Stack up
Not expecting high currents or any high-speed signals